TITLE:	METH030_Miseq_LIBRARY_PREP

1	INTRODUCTION

1.1	PURPOSE OF SOFTWARE

Preparing DNA samples for sequencing on the MiSeq NGS platform, using PCR, E-Gel electrophoresis, MN bead cleanup, QIAxcel measurement, QIAgility equimolar pooling and Bioanalyzer measurement.

1.2	DEFINITIONS AND ABBREVIATIONS

-

1.3	STATUS

The status and management of this document are regulated in the AODocs platform, the foundation wide accessibility is provided on N=Info.

2	RESPONSIBILITIES & PEOPLE

2.1 	RESPONSIBILITIES

| FUNCTION | ROLES and RESPONSIBILITIES |
| --- | --- |
| Lab workers (Researchers, Students, Guests) | Is able to use this document if contributing to the DNA barcoding work process |
| (Senior) Research Technicians | Knows these document instructions and is able to instruct other lab workers who need to use the Geneious program. |
| Head Laboratories | Has final responsibility for all laboratory workflow and instructions |

2.2	LIST OF PEOPLE ADDRESSED

Elza Duijm: 1st responsible for NGS post lab, Ion Torrent, QIAxcel and QIAgility

Marcel Eurlings: 2nd responsible for NGS post lab, Ion Torrent, QIAxcel and QIAgility

3	ADDITIONAL AIDS AND RESOURCES

4 	INSTRUCTIONS

4.1 	1st PCR DNA TEMPLATE

DNA template has to be amplified with a 20μl PCR using Nextera-tailed primers. A PCR sheet template, containing reaction mixes, primers and PCR programs is available in

B:\DNA Barcoding\Templates\PCR_template_NXT-1_20ul_v1.4

Do not use more then 35 cycles for this first round PCR!

Sequences Nextera tails:

Forward: 5'-TCGTCGGCAGCGTCAGATGTGTATAAGAGACAG-3'

Reverse: 5'-GTCTCGTGGGCTCGGAGATGTGTATAAGAGACAG-3'

4.2 	E-Gel

PCR products have to be checked on an E-Gel, using working instruction WI_SP0004_E-GEL

Make sure you use filter tips for pipetting from your PCR sample.

4.3	MN BEADS CLEANUP OF NXT PCR-1

PCR products of NXT PCR1 have to be cleaned before they can be labeled with sample-unique lables in NXT PCR2.

4.3.1	MN beads cleanup of 96-well plate, using magnetic extractor stamp

- Fill a 96- well round bottom plate with appropriate amount of magnetic beads

(in case of a standard clean-up 0.9 times the PCR volume)

- Add the PCR product to the magnetic beads, mix and incubate 5 minutes

- Fill 2 fresh round bottom plates with 50µl 80% EtOH (wash plates)

- Fill 1 fresh round bottom plate with appropriate (usually the volume of PCR product) amount of MQ (elution plate)

- Load the magnetic stamp with a cover plate and make sure the plate is locked into place

- Place the magnetic stamp in the sample / beads plate and wait for 1 minute

- Bring the stamp to the first wash plate and leave for 30 seconds

- Bring the stamp to the second wash plate and leave for 30 seconds

- Take out the stamp from the wash plate and let the beads dry for 30 seconds by holding it in the air

- Place the stamp in the elution plate and release the cover plate in a very quick snapping motion to  prevent beads from following the magnetic pins up the sides of the cover plate.

- Swirl the cover plate in the MilliQ to release all beads, then put the cover plate back on the stamp in the same position

- Mix the beads by pipetting the solution up and down to maximize the release of DNA from the beads

- Place the stamp back into the cover plate and wait for 5 minutes

- Take out the stamp with the cover plate and discard the cover plate (with attached beads)

- Transfer the purified PCR product to a 0.2 ml PCR plate for storage

4.3.2	MN beads cleanup of 96-well plate, using tips (optional)

- Let the MN beads get to room temperature (remove from fridge 15 min before use)

- Prepare fresh 80% EtOH (200µL per sample, e.g. 41,7 ml EtOH 96% + 8,3 ml MQ)

- Mix the MN beads well by vortexing

- Add 0.9x (PCR product volume) MN beads to the wells of a new 96-well plate

- Add the PCR product to the MN beads and mix well by pipetting

- Incubate at room temperature for 5 minutes

- Place the plate on the magnetic rack

- When the solution is clear (5 min), carefully remove the supernatant (LEAVE THE PLATE ON THE MAGNET)

- Wash the beads 2x with 100µl 80% ethanol (LEAVE THE PLATE ON THE MAGNET)

- Let the beads air dry for 1 minute (LEAVE THE PLATE ON THE MAGNET)

- Take the plate off the magnet and resuspend the beads with e.g. 25µl MQ, by pipetting up and down

- Place the plate back on the magnet

- When the solution is clear (5 min), move 22µl of the the supernatant to a clean plate

4.4	2nd PCR NEXTERA XT LABELS

PCR products from the 1st PCR have to be labeled with MiSeq Nextera XT labels. 384 unique label combinations are possible. A PCR sheet template, containing reaction mixes, primers and PCR programs is available in B:\DNA Barcoding\Templates\PCR_template_NXT-2_20ul_v1.5.xlsx

Do not switch well positions within your plate (e.g. when there are empty positions) and always use a multichannel pipette to avoid confusion.

The following instructions have to be followed when preparing the Nextera XT PCR:

- Prepare master mix: clean room

- Use fresh pipette tips when adding primers from 96-well plate

- Add template: NGS post lab

4.5	MEASURING 2nd PCR ON QIAXCEL

4.5.1 	Sample prep

- Fill 1 well with 1:3 dilution Analysis Marker (ladder)

- Fill sample wells with 18μl MQ + 2μl 2nd PCR product

- Fill empty wells with 10μl MQ

4.5.2 	Set parameters

- Open QIAxcel ScreenGel,  user: QIAxcel pass: QIAxcel2

- \Process\Process Profile\Library_prep

- \Sample Selection\Plate ID

- Adjust plate name

- \Sample Selection\Sample Row Selection (right window)

- Highlight relevant (all) rows by left clicking on them

- Right click on the well containing the Analysis Marker (ladder) and select \Toggle Analysis Marker

- \Run Check\Please Confirm\☑ All selected sample rows contain samples\☑ Alignment marker is loaded\☑ Size marker is loaded\Run

4.5.3	Analysis

- Tab \Analysis\

- After the run (80 minutes for 96 samples), select correct Plate ID, select all samples (Ctrl+A) and drag them into the right window

- Tab \Electropherogram\

- Left click on the Analysis Marker and switch to \Show single electropherogram view (button top left)

- \Analysis Properties (right window) \Default DNA\Reference Marker\No Marker\Start Analysis

- Tab \Reference Marker\Apply

- Tab \Electropherogram\Select All (button top left)\Analysis Properties\Lib_prep_smear\Start Analysis

- Switch to \Show single electropherogram view and check/adjust the smear of each well

- False peaks can be removed by selecting them with left click, followed by right click \Delete selected peak(s)

4.5.4	Report

- \Show electropherogram overview (button top left)\Select All

- Tab \Report (right window)\☑ PDF\Start Report/Export

- Uncheck \PDF\☑ XML Export\Start Report/Export

4.6	QIAGILITY EQUIMOLAR POOLING

4.6.1	Preparation QIAgility import

- Open the XML Export e.g. C170612A29_2017-09-04_11-29-42_20170904_133105_Ex.xml with Excel

- Copy columns T (ns1:SamplePosition), AZ (ns1:MedianSizeAreaOfInterest) and BG (ns1:MolarityAreaOfInterest) to a new tab

- Select the 3 columns

- \Gegevens\Duplicaten verwijderen\Ok\Ok

- Check if there are 96 samples (97 rows) present

- \Filter

- Check if the highest value divided by the lowest value ≤ 5 \Ok

- If the above is not the case, adjust the lowest values to highest value divided by 5 (e.g. if the highest value = 2.00, adjust all values ≤ 0.4 to 0.4)

- Fill in new lowest value (in this case 0.4) in wells with no value at all

For example:	ns1:MolarityAreaOfInterest

0.27 	-> 	0.4

0.0	->	0.4

- \Save as\Tekst (tab is scheidingsteken) (*.txt)\Ok\Ja

- Close Excel without saving changes to the original XML Export

4.6.2	QIAgility import

When the QIAgility robot cannot communicate with the PC, switch off the QIAgility, unplug and replug the QIAgility USB from the PC and restart QIAgility.

- Switch on the QIAgility liquid handling workstation before opening the QIAgility software

- Open QIAgility software

- Tab \Recent\normalize and pool__\Open

- Highlight area C1 by left clicking on it

- \Sample Data (right window)\Clear

- \Sample Data (right window)\Import...

- Select .txt import \Import\Finish

4.6.2.1	Delete empty wells (optional)

- Empty wells can be removed by right-clicking on area C1

- \Remove selected wells from sample bank...

- In the window Update Sample Bank, select wells to be removed by left-clicking on them while holding down Ctrl

- In the field Existing Bank, select bank1 (.. wells)

- Click button Delete Selection and Close

4.6.3	Sample normalization

- Highlight area C2 by left-clicking on it

- \Reaction List (top right)

- Double-click on Normalize samples of bank1 to conc...

- In the Sample Normalization window, adjust Final volume to the amount of available PCR product. If you have < 15μl left after the QIAxcel measurement(s), add 3μl to this number.

- Adjust Desired concentration to the Lowest non-zero sample conc (in this case 0.40)

- Press Tab to apply

- The NOTES: field should be empty

4.6.4	Refill tips and MQ

- Wearing gloves, remove empty tip racks from the QIAgility and replace them with full racks.

- When placing full racks in the QIAgility frame, make sure to align them correctly, so the notch in the plate aligns with the white dot on the frame

- Ensure the new tip rack closely connects with the frame by pressing along the ridges with 2 hands

- Replace the diluent and make sure the reservoir is half-full with MQ

- Select de wells in the refilled areas by dragging a field around them and right clicking

- \Set selected tips to ‘Available’

- Place 4 new Eppendorf low-bind 1.5ml eppies in the marked positions in area R1

- Place 2nd PCR plate in area C1 (Standard is Eppendorf twin.tec PCR plate 96 skirted. If you have a different plate type, highlight area C1 by left-clicking on it\right-click\Change plate type\Ok\)

- Place empty skirted PCR plate in area C2

- Click Start button

- Check all boxes and click Run

4.7	MN BEAD CLEANUP OF POOL

- Let the MN beads get to room temperature (remove from fridge 15 min before use)

- Prepare fresh 80% EtOH (1ml per sample)

- Pool the contents of the 4 eppies together and measure the total volume by pipetting

- Mix the MN beads well by vortexing

- Add 0.9x (pool volume) MN beads to the pool and vortex

- Incubate at room temperature for 5 minutes

- Place the plate on the magnetic eppie rack

- When the solution is clear, carefully remove the supernatant (LEAVE THE PLATE ON THE MAGNET)

- Wash the beads 2x with 500µl 80% ethanol (LEAVE THE PLATE ON THE MAGNET)

- Let the beads air dry for 1 minute (LEAVE THE PLATE ON THE MAGNET)

- Take the plate off the magnet and resuspend the beads with 300µl MQ, by pipetting up and down

- Place the plate back on the magnet

- When the solution is clear, move 280µl of the the supernatant to a clean low-bind eppie

4.8	MEASURE  SUBPOOLS ON QIAXCEL

-Measure the subpools in triplo on the QIAxcel

- Use the average molar concentration of the three measured subpools for pooling of these pools

4.8.1 Pooling of pools

If equimolar pooling of different pools is required, follow this calculation to dilute the pools to the lowest concentration.

example: 	pool Ⅰ   = 100 nM

pool Ⅱ  = 120 nM

pool Ⅲ = 250 nM

Adding an equal volume of the three pools together will give the right end concentration of 100 nM if pool Ⅱ and pool Ⅲ are also diluted to 100 nM.

Diluting pool Ⅱ to 100 nM = 120:100=1.2 So dilute pool Ⅱ 1.2 times (e.g. = 50µl pool Ⅱ + 10 µl MQ)

Diluting pool Ⅲ to 100 nM = 250:100=2.5 So dilute pool Ⅲ 2.5 times (e.g. = 50µl pool Ⅲ + 75 µl MQ)

Add an equal volume of pool Ⅰ, diluted pool Ⅱ and diluted pool Ⅲ together to create the 100 nM end pool.

4.8.2 Pooling of pools with a different number of samples

example: 	pool Ⅰ:   50 samples = 100 nM

pool Ⅱ:  10 samples = 120 nM

pool Ⅲ: 20 samples = 250 nM

First adjust the molarity of the pools to the lowest molarity (100nM in this example):

example:	pool I:   Already	 100nM

pool II:  120nM / 100nM = 1.2x dilution eg. 10ul MQ + 50µl pool II

pool III: 250nM / 100nM = 2.5x dilution eg. 75ul MQ + 50µl pool III

Now the molarity of the 3 pools is the same (100nM). To make the final pool of these 3 pools, you have to correct for the number of samples in each pool. To do this, pipet the amount of µl corresponding to the number of samples in that pool into a new tube. For this example:

pool I: 	50µl

pool II: 	10µl

pool III:	20µl

4.8.3 Measure endpool on BIOANALYZER

Measure the end pool in triplo in the following dilutions: 3x 1:10 .  3x 1:100 and  3x undiluted

- Let the High Sensitivity Kit chemicals get to room temperature (remove from refrigerator 20 min before use)

- Fill cleaning chip with 350µl MQ and place in Bioanalyzer

- Prepare a High Sensitivity Chip, according to the manual, which is available in the lab and can be downloaded here: G2938-90321_SensitivityDNA_KG_EN

- Remove cleaning chip from Bioanalyzer

- Start Bioanalyzer software, indicate number of samples and provide sample names

- Place the High Sensitivity Chip in the Bioanalyzer

- Start run

- After run, select each sample by left-clicking on them (top-left)

- Tab \Region Table\ and adjust smear (blue lines) if necessary

- Write down the Molarity in pM [=pmol/l] for each sample of the same dilution and calculate the average values of the 1:10 and undiluted pool.

- Calculate the molarity of the pools in nM (in this example is the picture from the 10 times diluted sample meaning the concentration is 26.80 nM).

Safe the report as a pdf file.Baseclear needs 20ul of a minimum 5nM endpool

5	ADDITIONAL PROCEDURES

5.1	For safety and lab regulations please reads the following documnets:

PROC002_GENERAL_LAB_REGULATIONS

PROC003_SAFETY_PROCEDURES

WI_WS001_PRE-LAB

WI_WS_NGS_LAB

6	CALCULATIONS

-

7	QUALITY CONTROL

-

8	RELATED/ASSOCIATED DOCUMENTS

-

9	LITERATURE

-